## Sheet1
Current text | Proposed text | Comment | Summary | TGA Comment | Action | Resolved?
Risk proportionate regulation |  | In response to the 2015 Expert Review of Medicines and Medical Devices Regulation, the TGA reformed the regulation of a number of low-risk products, including disinfectants. Due to their low-risk nature, disinfectants were downregulated, the extent dependant on the claims that were being made. For example, disinfectants not making specific claims shifted from being 'listed' to being exempt from inclusion in the ARTG, and disinfectants making specific claims shifted from 'registered' to 'listed'. While these categorisations are still in place, the requirements for these products do not greatly differ. Respondent encourages the TGA to ensure that disinfectant regulations are risk proportionate and fit for purpose, considering the statement that these products are low-risk. | Request to review disinfectant regulation, based on risk categorisation of products |  | Refer to MDAB for further consideration | Yes
International alignment |  | Respondent strongly believes that the acceptance criteria should be aligned with that accepted by comparable overseas regulators, or at a minimum, maintaining the current Australian specific criteria with the option for TGA to accept data/ decisions from comparable overseas regulatory authorities. Any extra testing causes Sponsors significant cost and time delays which we hope you will take into great consideration. | Request the acceptance of overseas regulatory decisions | A review of international approaches to disinfectant regulation was performed, with obvious differences that are not easily overcome. | Rejected | Yes
International alignment |  | Respondent requests an increase in international alignment within the Instructions. This will reduce costs to industry and consumers without increasing risks. We are not aware of any benefits of unique Australian requirements. Respondent encourages the TGA to consider international alignment in all facets of its work, including the Instructions. The Instructions are unique to Australia and often require sponsors to conduct further microbiological testing for products already on the market in other jurisdictions. The extra testing comes at a significant cost and time delays with no identifiable benefits. In many cases, due to the size of the Australian market compared to these costs, internationally available products are not brought to market here. A lesson learnt during the COVID pandemic is that having a robust and varied supply chain is critical to withstand significant global events. Reducing supply without established and considered benefits does not serve the Australian public. | Request greater alignment with international requirements | A review of international approaches to disinfectant regulation was performed, with obvious differences that are not easily overcome. However, this update is seeking to reduce the differences where possible. | Rejected | Yes
Table of Contents |  | The current contents page of the Instructions does not detail any of the headings in Part 2 Division 2. The expansion of this section of the contents would allow for easier navigation to the relevant parts of the document. | Request to tease out contents page |  | More detail added to contents page | Yes
p4 Introduction Disinfectants and cleaners intended for use on medical devices are regulated as medical devices under Chapter 4 of the Act, and are required to comply with the essential principles identified in Schedule 1 of the Therapeutic Goods (Medical Devices) Regulations 2002 |  |  | Respondent is concerned about lack of guidance related to medical device disinfectant products. Historical guidance (still being used as a reference) does not mention wipes or spray products. Respondent requests further guidance about other medical device disinfection systems, such as UVC cabinets. | Out of scope. For note by MDAB. Note that MD wipes and sprays are evaluated with reference to both the Instructions and the historical guidelines. | No action required. Refer to MDAB. | Yes
p5 Part 1 - The TGA disinfectant test |  |  | To move Part 1 – The TGA disinfectant test to sit as an appendix to improve readability of the document. | Renumbering or rearrangement in Instructions would require major rewrite of TGO 104. TGO 104 requires only minor updates | Rejected | Yes
p5 2. The test consists of the following steps, undertaken in accordance with the instructions or requirements in sections 2 – 8 of the TGA Disinfectant Test and in the following order: |  |  | Text states that it is referring to sections 2-8 of the TGA Disinfectant Test, but this is unclear. |  | Amended to read "sections 2-8 of the TGA Disinfectant Test (Part 1 of this document)". Numbers added to headings in Part 1. | Yes
p5 TGA test organisms |  | Consider aligning these organisms with those in EN 13727 | Propose aligning organisms across different test methods. | The TGA Disinfectant Test method was developed and verified using the stated organisms. Changes to the test organisms would require studies to verify and rewrite the method. The TGA Disinfectant Test is a stand alone test. The Instructions state that other test methods can be used. | Rejected | Yes
|  | The instructions state that EN 13697 can be used as acceptable test method for bactericidal testing and/or fungicidal testing. We would like to propose that testing against yeast and fungi are not required if only making a bactericidal claim and that this be clearly written in the Instructions. | Propose adapting test methods to use only the relevant organisms for the claims being made. Specifically for EN 13697 | EN 13697 is for bactericidal AND/OR fungicidal efficacy testing, so only the organisms for which the claims are being made need to be tested. | No further action required | Yes
|  | Under virucidal testing, EN 14476 is accepted test method but only if Herpes Simplex is added to the organisms tested. The rationale for this is unclear and believe that the TGA’s rationale for this should be included in the Instructions. | Request reasoning for inclusion of Herpes Simplex to virucidal testing when using EN 14476 | Unable to provide this background (note that we haven't assessed EN 14476 using Herpes Simplex, only the specific claims, or the specified viruses from EN 14476 for a general virucidal claim). Remove Herpes Simplex from Instructions. | Herpes Simplex removed from Virucidal testing (p16) | Yes
Virucidal testing |  | To make the same claim the Instructions require a log 4 reduction in virus testing while the United States (US) Food and Drug Administration (FDA) only require a log 3. We would like to understand from the TGA the scientific basis of the higher pass rates in Australia compared to the US, for example, whether there are significant clinical differences that can be observed. | Request reasoning for difference in virucidal log reduction from US FDA | Unable to provide this background. Note that most virucidal test methods e.g EN 14476 and EN 16777 require a 4 Log reduction | Rejected | Yes
|  | Several test methods (EN 13727, EN 14561, EN 13697, EN 1276, and EN 13704) from the Instructions require three testing concentrations (e.g. including one in the active range and one in the non-active range) so that formulations can be better understood during development phase. During pre-consultation TGA agreed that when three concentrations were specified in the test method, they would accept one test concentration in the active range. The Consultation also have additional statement "Testing should also reflect the in-use conditions of the product.” We do not believe that this is clear and that additional clarifying statements are required to be added to the Instructions. | Request further clarification that a single product concentration requires testing (rather than 3 concentrations) | This was previously agreed. Clarification added: Some of the test methods referenced in this document specify that testing should be performed at three product concentrations. This is applicable to the product development phase, and testing at the in-use concentration of the product only is expected once a product is out of development. | Clarification added | Yes
|  | We also would like to propose to consider adding EN 17387 (European Standard EN 17387: Chemical disinfectants and antiseptics - Quantitative test for the evaluation of bactericidal and yeasticidal and/or fungicidal activity of chemical disinfectants in the medical area on nonporous surfaces without mechanical action - Test method and requirements (phase 2, step 2)) as acceptable bactericidal test method. | Request adding EN 17387 as an acceptable bactericidal test method | EN 17387 standard is for disinfectants used in the medical area is relatively new (2021) and is the replacement of EN 13697 - standard used in food, industrial, domestic and institutional areas. Differences in interfering substances and criteria. EN 13697 requires 4 Log reduction for bacteria and 3 Log for fungi whilst EN 17387 requires 5 Log reduction for bacteria and 4 Log reduction for fungi. Add as an acceptable method. | Added references to EN 17387 as appropriate | Yes
p9 TGA test procedure 6.3 At 8 minutes, subculture one drop (0.02mL ± 0.002mL) into each of 5 tubes containing recovery broth. To ensure delivery of 0.02mL into the first tube of recovery broth at exactly 8 minutes, it will be necessary to withdraw a suitable amount from the disinfectant test mix shortly beforehand. This must be immediately preceded by vortexing. Surplus sample must be returned to the test mix. | p9 TGA test procedure At the proposed product contact time as per product’s usage instruction, subculture one drop (0.02mL ± 0.002mL) into each of 5 tubes containing recovery broth. To ensure delivery of 0.02mL into the first tube of recovery broth at exactly 8 minutes, it will be necessary to withdraw a suitable amount from the disinfectant test mix shortly beforehand. This must be immediately preceded by vortexing. Surplus sample must be returned to the test mix. |  | Request allowing modification of the TGA Disinfectant test method to reduce the contact time 'as per product usage instructions' | The TGA Disinfectant Test should be performed as written with an 8 minute test period, and the claimed contact time should be confirmed using other test methods. Although a reduction in the test time has a reasonable scientific basis, it is not clear to what extent changes to the test would invalidate the test method, due to the inclusion in the test method of a rechallenge. The option exists to use other EN tests that allow for changes to the contact time. | Rejected | Yes
p9 TGA test procedure 6.3 At 8 minutes, subculture one drop (0.02mL ± 0.002mL) into each of 5 tubes containing recovery broth. To ensure delivery of 0.02mL into the first tube of recovery broth at exactly 8 minutes, it will be necessary to withdraw a suitable amount from the disinfectant test mix shortly beforehand. This must be immediately preceded by vortexing. Surplus sample must be returned to the test mix. |  | Identification of issue • Contact time for the option A, B and C tests is specified to be 8 minutes. (Part 1 Section 6 of the Instructions) • It is logical to alter the contact time depending on the products’ directions for use. For example, if the directions for use stipulate a 5-minute contact time then this time should be used while completing the testing. • This logical variation in test method does not create any performance or efficacy issues. Suggested solution • Rewrite the instructions to allow the contact time for the option A, B and C tests to be varied to align with product directions for use. | Request allowing modification of the TGA Disinfectant test method to reduce the contact time | The TGA Disinfectant Test should be performed as written with an 8 minute test period, and the claimed contact time should be confirmed using other test methods. Although a reduction in the test time has a reasonable scientific basis, it is not clear to what extent changes to the test would invalidate the test method, due to the inclusion in the test method of a rechallenge. The option exists to use other EN tests that allow for changes to the contact time. | Rejected | Yes
p11 Part 2 - Disinfectant performance tests |  | Check Part for alignment of terminology with TGO 104 |  |  | Checked | Yes
p11 3 Antimicrobial stability A disinfectant must pass at least one suitably sensitive test according to the appropriate evaluation guidelines for the grade of disinfectant (see Table 2 below) at the end point at which testing is occurring, with the levels of active(s) at the predicted final level. | p11 3 Antimicrobial stability Efficacy testing is required to account for such effects as interactions with the container, effects of storage, transport conditions, efficacy under usage conditions such as the use of the product on surfaces or use at the intended dilution. If the active(s) is stable according to chemical and physical stability data, then there is no need for microbial efficacy testing at the end of shelf-life. However, if the active(s) concentration decreases by more than 10% during the shelf life of the biocidal product, then efficacy tests should be performed at the endpoint to confirm the shelf life. If end of shelf-life efficacy is required then a disinfectant must pass at least one suitably sensitive test according to the appropriate evaluation guidelines for the grade of disinfectant (see Table 2 below) at the end point at which testing is occurring, with the levels of active(s) at the predicted final level. |  | Request that microbiological stability testing not be required, to align internationally, and with the APVMA. | As discussed in previous meetings, microbiological stability testing is required due to the possible interactions between the product disinfectant and the packaging. | Rejected | Yes
p11 3 Antimicrobial stability A disinfectant must pass at least one suitably sensitive test according to the appropriate evaluation guidelines for the grade of disinfectant (see Table 2 below) at the end point at which testing is occurring, with the levels of active(s) at the predicted final level. |  |  | Propose that prior microbiological testing on a “worst case product” at the end of shelf-life, plus routine chemical and physical testing, is sufficient to confirm stability. Worst-case product would be in alignment with the BPR definition (a decrease in active content of up to 10%). | As discussed in previous meetings, microbiological stability testing is required due to the possible interactions between the product disinfectant and the packaging. In the absence of appropriate supporting evidence, the TGA can not currently accept this proposal. This may be reconsidered in future. | Rejected | Yes
p14 Division 2 - Microbial efficacy general requirements Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable . |  | Does this mean that the TGA will now also accept the log reductions included in those standards as ‘pass’ criteria. If so, this should be stated. If not, where are the pass criteria to be met stated? They are not in TGO 104. | Are acceptance criteria for referenced test methods used to determine compliance? | Generally yes, with some exceptions, which are specified in the Instructions. | No further action required (though document checked for clarity) | Yes
p14 Division 2 - Microbial efficacy general requirements Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable . |  | We would like to propose that modification of test methods for the relevant area of application, with scientific justification, be allowed, such that they reflect the in-use conditions of the products and the appropriate organisms. We also proposed to have an option when using international test method, that other organisms that are representative of the mandatory organisms (gram-positive, gram-negative, enveloped and non-enveloped viruses) would be allowed. | Request allowing modification of test methods to reduce the contact time or use alternative organisms | Most EN methods allow for these changes, within limits. Changes are permitted and assessed within the ranges written into the method. General claims need to be substantiated by the organisms stated in the test method, but additional specific claims can be added. | No further action required | Yes
p14 Division 2 - Microbial efficacy general requirements Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable . |  | The instruction states that EN 13697 can be used as acceptable test method for bactericidal testing and/or fungicidal testing. We would like to propose that testing against yeast and fungi are not required if only making a bactericidal claim and that this be clearly written in the Instructions. | Request allowing use of only certain organisms in test method (specifically EN 13697), rather than full suite of organisms | EN 13697 is for bactericidal AND/OR fungicidal efficacy testing, so only the organisms for which the claims are being made need to be tested. General claims need to be substantiated by the organisms stated in the test method, but additional specific claims can be added. | No further action required | Yes
p14 Division 2 - Microbial efficacy general requirements Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable . |  | Identification of issue • The Instructions state that EN 13697 can be used as a test method for bactericidal testing and/or fungicidal testing. • Not all products that use EN 13697 as a test method make claims for yeast and fungi. • It is our understanding that EN 13697 is designed in modules to allow the testing required to be completed rather than having to complete the entire test. o The Scope of EN 13697 states that “this European Standard specifies a test method (phase 2/step 2) and the minimum requirements for bactericidal and/or fungicidal or yeasticidal activity of chemical disinfectants”. Table 1 – Experimental conditions of the standard is split into three different areas: bactericidal activity on non-porous surfaces without mechanical action, fungicidal activity on non-porous surfaces without mechanical action and yeasticidal activity on non-porous surfaces without mechanical action. o Appendix 4 Overview of standards, test conditions and pass criteria (PT 1-5) of the Guidance on the Biocidal Products Regulation (BPR) show bacteria, yeast and fungal for EN 13697 as separate. Both yeast and fungal microorganisms testing being shown as optional in some cases. • It is not clear in the Instructions if EN 13697 can be varied in line with the product claims they are making. o For example, if no yeast or fungicidal claims are being made, only the bacteria testing from EN 13697 would be required. Suggested solution • Rewrite the Instructions to allow the variation of EN 13697. | Request allowing use of only certain organisms in test method (specifically EN 13697), rather than full suite of organisms | EN 13697 is for bactericidal AND/OR fungicidal efficacy testing, so only the organisms for which the claims are being made need to be tested. General claims need to be substantiated by the organisms stated in the test method, but additional specific claims can be added. | No further action required | Yes
p14 Division 2 - Microbial efficacy general requirements Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable . |  | Under virucidal testing, EN 14476 is accepted test method but only if Herpes Simplex is added to the organisms tested. The rationale for this is unclear and believe that the TGA’s rationale for this should be included in the Instructions. | Request reasoning for inclusion of Herpes Simplex to virucidal testing when using EN 14476 | Unable to provide this background (note that we haven't assessed EN 14476 using Herpes Simplex, only the specific claims, or the specified viruses from EN 14476 for a general virucidal claim). Remove Herpes Simplex from Instructions. | Herpes Simplex removed from Virucidal testing | Yes
p14 Division 2 - Microbial efficacy general requirements Testing should be performed according to the test method used with regard to test organisms, number of carriers and test concentrations tested, and type and concentration of organic soil to be used to simulate dirty and clean conditions. Testing should also reflect the in-use conditions of the product. |  | Several test methods (EN 13727, EN 14561, EN 13697, EN 1276, and EN 13704) from the Instructions require three testing concentrations (e.g. including one in the active range and one in the non-active range) so that formulations can be better understood during development phase. During pre-consultation TGA agreed that when three concentrations were specified in the test method, they would accept one test concentration in the active range. The Consultation also have additional statement "Testing should also reflect the in-use conditions of the product.” We do not believe that this is clear and that additional clarifying statements are required to be added to the Instructions. | Request further clarification that a single product concentration requires testing (rather than 3 concentrations) | This was previously agreed. Clarification added: Some of the test methods referenced in this document specify that testing should be performed at three product concentrations. This is applicable to the product development phase, and testing at the in-use concentration of the product only is expected once a product is out of development. | Clarification added | Yes
p14 Division 2 - Microbial efficacy general requirements Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable . |  | We would like to propose to consider adding EN 17387 (European Standard EN 17387: Chemical disinfectants and antiseptics - Quantitative test for the evaluation of bactericidal and yeasticidal and/or fungicidal activity of chemical disinfectants in the medical area on nonporous surfaces without mechanical action - Test method and requirements (phase 2, step 2)) as acceptable bactericidal test method. | Request adding EN 17387 as an acceptable bactericidal test method | EN 17387 standard is for disinfectants used in the medical area is relatively new (2021) and is the replacement of EN 13697 - standard used in food, industrial, domestic and institutional areas. Differences in interfering substances and criteria. EN 13697 requires 4 Log reduction for bacteria and 3 Log for fungi whilst EN 17387 requires 5 Log reduction for bacteria and 4 Log reduction for fungi. Add as an acceptable method. | Added references to EN 17387 as appropriate | Yes
p14 Division 2 - Microbial efficacy general requirements Testing should be performed according to the test method used with regard to test organisms, number of carriers and test concentrations tested, and type and concentration of organic soil to be used to simulate dirty and clean conditions. Testing should also reflect the in-use conditions of the product. |  |  | Is virucidal testing perormed as per recognised test method? A 4log reduction may not be achievable in that case. | It is achievable. Note that most virucidal test methods e.g EN 14476 and EN 16777 require a 4 Log reduction | Rejected | Yes
p14 Division 2 - Microbial efficacy general requirements Testing should be performed according to the test method used with regard to test organisms, number of carriers and test concentrations tested, and type and concentration of organic soil to be used to simulate dirty and clean conditions. Testing should also reflect the in-use conditions of the product. | p14 Division 2 - Microbial efficacy general requirements Testing should be performed according to the test method used with regard to test organisms, number of carriers and test concentrations tested, and type and concentration of organic soil to be used to simulate dirty and clean conditions. Testing should also reflect the in-use conditions of the product. Modification of test method for relevance to the area of application with scientific justification must be allowed to reflect the in-use conditions of the product & the appropriate organisms. |  | Request allowing use of only certain organisms in test method (specifically EN 13697), rather than full suite of organisms | EN 13697 is for bactericidal AND/OR fungicidal efficacy testing, so only the organisms for which the claims are being made need to be tested. | No further action required | Yes
p14 Division 2 - Microbial efficacy general requirements Testing should be performed according to the test method used with regard to test organisms, number of carriers and test concentrations tested, and type and concentration of organic soil to be used to simulate dirty and clean conditions. Testing should also reflect the in-use conditions of the product. | p14 Division 2 - Microbial efficacy general requirements Testing should reflect the in-use conditions of the product such as contact time and disinfectant concentration. Testing should be performed according to the test method used with regard to test organisms, number of carriers, and type and concentration of organic soil to be used to simulate dirty and clean conditions. |  | Request allowing modification of test methods to reduce the contact time or use alternative organisms | Most EN methods allow for these changes, within limits. Changes are permitted and assessed within the ranges written into the method. | No further action required | Yes
p14 Division 2 - Microbial efficacy general requirements Regardless of whether the product is used undiluted or diluted by the user to the Minimum Recommended Concentration, the product should be formulated at the lower extreme of the active ingredient in the product specification, stored to expiration and then subjected to the microbial efficacy tests, and this data should be available on request. | p14 Division 2 - Microbial efficacy general requirements Regardless of whether the product is used undiluted or diluted by the user to the Minimum Recommended Concentration, the product should be formulated at the lower extreme of the active ingredient in the product specification, stored to expiration and then subjected to the microbial efficacy tests, when required. This data should also be made available on request. |  | Request that microbiological stability testing not be required. | As discussed in previous meetings, microbiological stability testing is required due to the possible interactions between the product disinfectant and the packaging. | Rejected | Yes
p14 Division 2 - Microbial efficacy general requirements The product may be tested before the end of the shelf life to initially provide a full set of data for evaluation purposes if product stored to the end of the expiration period is not available. The data can be generated from a single batch of product. However, data from further batches may be requested by the TGA in certain circumstances (i.e., the product contains a novel active or the single batch stability data raised concerns) According to the microbiological stability requirements specified in Section 4 above, once product stored to the end of the expiration period becomes available. End of shelf life data should be generated from a minimum of two batches. | p14 Division 2 - Microbial efficacy general requirements The product may be tested before the end of the shelf life to initially provide a full set of data for evaluation purposes if product stored to the end of the expiration period is not available. The data can be generated from a single batch of product. However, data from further batches may be requested by the TGA in certain circumstances (i.e., the product contains a novel active or the single batch stability data raised concerns). If the active(s) concentration decreases by more than 10% during the shelf life of the biocidal product, then efficacy tests should be performed at the endpoint to confirm the shelf life according to the microbiological stability requirements specified in Section 4 above, once product stored to the end of the expiration period becomes available. End of shelf life data should be generated from a minimum of two batches. |  | Request that microbiological stability testing not be required unless chemical or physical testing does not comply | As discussed in previous meetings, microbiological stability testing is required due to the possible interactions between the product disinfectant and the packaging. | Rejected | Yes
p14 Division 2 - Microbial efficacy general requirements Test methods must be validated by individual laboratories for each test method used in accordance with tests which have been validated or refereed at national or international level. This should involve validation of individual operators by performance of the specified test using a control or product that is known to pass the specified test at a given concentration as well as testing to determine the most appropriate neutraliser/inactivator, if necessary. | p15 Division 2 - Microbial efficacy general requirements Test methods must be validated by individual laboratories for each test method used in accordance with tests which have been validated or refereed at national or international level. This should involve validation of individual operators by performance of the specified test using a control or product that is known to pass the specified test at a given concentration as well as testing to determine the most appropriate neutraliser/inactivator, if necessary. The most stringent claim is regarded as the organism least susceptible to the test product. In general, the hierarchy of susceptibility from hardest to easiest to kill is bacterial spores, mycobacteria, non-enveloped viruses, fungi, vegetative bacteria and enveloped viruses. There are other formulation considerations such as pH and incompatible excipients/actives. |  | Propose moving hierarchy of kill information to be clearly applicable to all disinfectants (from wipes section) | Move hierarchy of kill information | Hierarchy of kill information moved to General Requirements | Yes
p14 Division 2 - Microbial efficacy general requirements |  |  | Propose to add requirements summary table to the document for ease of use | Although it is acknowledged that tables can be useful, it is decided that, in this case, a summary table will not be able to contain all the information required, and is likely then to cause further confusion, while discouraging users of the document from reading the full document. | Rejected | Yes
|  | We suggest that a summary table is added so that the possible acceptable test methods and testing for each type of product could be quickly ascertained. This could be included in each section of the Instructions or as an appendix at the end of the document. An example of a table that clearly articulates requirements can be found in Understanding EN 14885 – Surface Disinfectants (including the surface of medical devices) in the Medical Area Test Standards. This could be used as inspiration for tables included in the Instructions | Propose to add summary table for ease of use | Although it is acknowledged that tables can be useful, it is decided that, in this case, a summary table will not be able to contain all the information required, and is likely then to cause further confusion, while discouraging users of the document from reading the full document. | Rejected | Yes
|  | Identification of issue • Related to the issue above, contact times in many methods are stipulated and may not be open to variation in the wording of the method. o E.g. The AOAC Hard Surface Carrier Test methods (991.47-49) specifies a contact time of 10 minutes. o E.g. EN 14561 is also quite prescriptive with a mandatory test time of 60 minutes and the option to also test at 5, 15 and/or 30 minutes. • It is logical to alter the contact time depending on the products’ directions for use. o For example, if the directions for use stipulate a 5-minute contact time then this time should be used while completing the testing. • This logical variation in test method does not create any performance or efficacy issues. Suggested solution • Rewrite the instructions to allow the contact time for methods to align with the directions for use. o E.g. This could be achieved by extending the current statement at the top of page 15 so that it reads, "All testing should be performed using the in-use product conditions including contact time(s), test methods may be modified if necessary to the extent necessary to accommodate contact time variation". | Request allowing modification of test methods to reduce the contact time | Most EN methods allow for these changes, within limits. Changes are permitted and assessed within the ranges written into the method. | No further action required | Yes
p15 Hospital grade disinfectants Bactericidal testing If the AOAC Use Dilution Test is chosen, the disinfectant passes when there is no growth in 59 out of 60 carriers per organism. If the Hard Surface Carrier Test is used, the criteria for pass or fail are described in the test method. If 10 carriers are used for any additional claims, no carriers may show growth. |  |  | Request clarification of the numbers of carriers used for each claim/test method. Rearrange final sentence so its clear that the 10 carriers are for additional claims. | Changed to read: If the AOAC Use Dilution Test is chosen, the disinfectant passes when there is no growth in 59 out of 60 carriers per organism. If additional claims against organisms not included in the original test are to be made (e.g. ‘Kills E coli 0157’), the extra organism(s) can be tested using 10 carriers rather than 60 when using the AOAC test methods. None of the 10 carriers should show growth. Soil should be included at a minimum of 5% blood serum and inorganic soil such as AOAC hard water. If the Hard Surface Carrier Test is used, the criteria for pass or fail are described in the test method. If the ASTM or EN test methods are used, applicable test conditions are as described in the test method, however a minimum 5-log10 reduction must be achieved. | Clarification added | Yes
p15 Hospital grade disinfectants Bactericidal testing If any specific biocidal claims are made (i.e., virucidal, fungicidal, tuberculocidal, sporicidal, or other biocidal activity), the disinfectant must pass appropriate tests as specified below. All tests should be carried out using the exposure time, temperature and pH specified on the label. For products intended for use on surfaces that have not been pre-cleaned, 5% organic soil must be included. |  |  | Is soil used as per the referenced test method, or as per Instructions? | Add comment that test methods are to be employed as written, with any exceptions specified in the Instructions. | Comment added that test methods are to be employed as written, with any exceptions specified in the Instructions. | Yes
p15 Hospital grade disinfectants Bactericidal testing These disinfectants MUST also pass a bactericidal carrier test. AOAC , methodology (60 carriers per organism) or a quantitative carrier test (methods such as ASTM E2197 , ASTM E2111 , EN 14561 or EN 13697) may be used. | p16 Hospital grade disinfectants Bactericidal testing These disinfectants MUST also pass a bactericidal carrier test. AOAC , methodology (60 carriers per organism) or a quantitative carrier test (methods such as ASTM E2197 , ASTM E2111 , EN 14561 or EN 13697 or EN 17387*) may be used. * European Standard EN 17387: Chemical disinfectants and antiseptics - Quantitative test for the evaluation of bactericidal and yeasticidal and/or fungicidal activity of chemical disinfectants in the medical area on non-porous surfaces without mechanical action - Test method and requirements (phase 2, step 2) |  | Request adding EN 17387 as an acceptable bactericidal test method | EN 17387 standard is for disinfectants used in the medical area is relatively new (2021) and is the replacement of EN 13697 - standard used in food, industrial, domestic and institutional areas. Differences in interfering substances and criteria. EN 13697 requires 4 Log reduction for bacteria and 3 Log for fungi whilst EN 17387 requires 5 Log reduction for bacteria and 4 Log reduction for fungi. Add as an acceptable method. | Added references to EN 17387 as appropriate | Yes
p16 Hospital grade disinfectants Fungicidal testing For a fungicidal claim, any test of reasonable scientifically-based, peer reviewed methodology will be considered, although carrier test methodology is preferred. Acceptable tests include the AOAC Use Dilution Test, which can be modified for fungi, with 10 carriers for each of two batches of product, and acceptance criteria of all fungal spores on all carriers killed, the AOAC Fungicidal Test , which can be modified for a carrier test, and the ASTM E2197 test methodology, with acceptance criteria as for the modified AOAC Use Dilution Test. Tests conducted in accordance with EN 13624 or EN 13697 will also be accepted. |  |  | Request clarification of the numbers of carriers used for each claim/test method | 60/10 applies to AOAC. EN methods are as per method. | Clarification included that 60/10 applies to AOAC method | Yes
p16 Hospital grade disinfectants Virucidal testing Possible alternatives for Poliovirus include Hepatitis A, Feline Calicivirus or Murine Norovirus. However, any company choosing a virus other than Poliovirus or Parvovirus for virucidal testing should justify the use of an alternative virus . |  |  | Is a justification still required for use of the named surrogate viruses? | Remove requirement for justification when a recognised and nominated surrogate virus is used. Need to check for any further references in the document. | Removed requirement for justification for recognised surrogate viruses | Yes
p18 Hospital grade disinfectant wipes |  | Suggest this heading be changed to something like ‘Specific requirements for certain disinfectant formats’ or similar that covers all the formats talked about below. |  |  | Rejected | Yes
|  | The heading ‘Hospital grade disinfectant wipes’ should be downgraded a level so it is black and bold the same as the ones for the sponges and sprays. |  |  | Editorial | Yes
|  | Also, it needs to be made clear whether the requirements in these sections are instead of, or additional to, the requirements in the previous section headed ‘Hospital grade disinfectants’ |  |  | Sentence added | Yes
p19 Hospital grade disinfectant wipes 1. Pass a bactericidal carrier test in accordance with the European Standard EN 16615 or alternatively, a modified AOAC Germicidal Spray Test, with 60 carriers, tested by wiping the surface of the carriers with the saturated wipe and subculturing the carriers after the specified contact time. One wipe should be used for a minimum of 10 carriers. The performance criteria should be no growth from 59/60 carriers. |  |  | Request clarification of the numbers of carriers used for each claim/test method |  | Clarification added | Yes
p19 Hospital grade disinfectant wipes 1. Pass a bactericidal carrier test in accordance with the European Standard EN 16615 or alternatively, a modified AOAC Germicidal Spray Test, with 60 carriers, tested by wiping the surface of the carriers with the saturated wipe and subculturing the carriers after the specified contact time. One wipe should be used for a minimum of 10 carriers. The performance criteria should be no growth from 59/60 carriers. |  |  | Request clarification regarding use of suspension or carrier testing |  | Clarification added throughout document | Yes
p19 Hospital grade disinfectant wipes 2. If specific biocidal claims are made, all claims should be supported with data from a suspension or carrier test, with added soil as appropriate . If a suspension test is chosen, the test should be conducted on the product after extraction from the wipes. These tests should be carried out according to the requirements of hospital grade disinfectants as above. |  |  | Is soil used as per the referenced test method, or as per Instructions? | Add comment that test methods are to be employed as written, with any exceptions specified in the Instructions. | Comment added that test methods are to be employed as written, with any exceptions specified in the Instructions. | Yes
p19 Hospital grade disinfectant wipes 3. In addition to verifying all biocidal claims with a carrier or suspension test, the product must also pass a simulated in-use test (such as EN 16615) against the most stringent claim being made, to demonstrate that the efficacy of the disinfectant is not reduced when combined with a wipe. The design of the test should be based on a carrier/surface test and could involve wiping the wipe over a carrier/surface and culturing both the carrier and the liquid that has been expressed from the used wipe. Alternatively, the entire wipe could be cultured. |  |  | Request clarification regarding use of suspension or carrier testing |  | Clarification added throughout document | Yes
p19 Hospital grade disinfectant wipes The most stringent claim is regarded as the organism least susceptible to the test product. In general, the hierarchy of susceptibility from hardest to easiest to kill is: bacterial spores, mycobacteria, non-enveloped viruses, fungi, vegetative bacteria and enveloped viruses. | p19 Hospital grade disinfectant wipes Delete |  | Propose moving hierarchy of kill information to be clearly applicable to all disinfectants (from wipes section) |  | Hierarchy of kill information moved | Yes
p19 Hospital grade disinfectant surface sprays (1) Pass a bactericidal carrier test such as the AOAC Germicidal Spray Test. Sixty carriers per organism should be used, plus 10 carriers for each additional bactericidal claim. The performance criteria should be no growth from 59/60 carriers. Testing in accordance with EN 14561 or EN 13697 will be accepted (test to be chosen depends on intended use – for hospital grade disinfectants, EN 14561 is the most appropriate). |  |  | Request clarification of the numbers of carriers used for each claim/test method |  | Clarification added throughout document | Yes
p20 Household/commercial grade disinfectant wipes (1) Pass a bactericidal carrier test in accordance with the European Standard EN 16615 or alternatively, a modified AOAC Germicidal Spray Test, with 60 carriers, tested by wiping the surface of the carriers with the saturated wipe and subculturing the carriers after the specified contact time. One wipe should be used for a minimum of 10 carriers. The performance criteria should be no growth from 59/60 carriers. Organic soil is not necessary, but inorganic soil should be used, in the form of AOAC Hard Water or equivalent. |  |  | Request clarification of the numbers of carriers used for each claim/test method |  | Clarification added throughout document | Yes
p20 Household/commercial grade disinfectant wipes (1) Pass a bactericidal carrier test in accordance with the European Standard EN 16615 or alternatively, a modified AOAC Germicidal Spray Test, with 60 carriers, tested by wiping the surface of the carriers with the saturated wipe and subculturing the carriers after the specified contact time. One wipe should be used for a minimum of 10 carriers. The performance criteria should be no growth from 59/60 carriers. Organic soil is not necessary, but inorganic soil should be used, in the form of AOAC Hard Water or equivalent. |  |  | Request clarification regarding use of suspension or carrier testing |  | Clarification added throughout document | Yes
p20 Household/commercial grade disinfectant wipes (3) In addition to verifying all biocidal claims with a carrier or suspension test , the product must also pass a simulated in-use test (such as EN 16615) against the most stringent claim being made to demonstrate that the efficacy of the disinfectant is not reduced when combined with a wipe. The design of the test should be based on a carrier/surface test and could involve wiping the wipe over a carrier/surface and culturing both the carrier and the liquid that has been expressed from the used wipe. Alternatively, the entire wipe could be cultured. |  |  | Request clarification regarding use of suspension or carrier testing |  | Clarification added throughout document | Yes
p21 Household/commercial grade disinfectant wipes The most stringent claim is regarded as the organism least susceptible to the test product. In general, the hierarchy of susceptibility from hardest to easiest to kill is: bacterial spores, mycobacteria, non-enveloped viruses, fungi, vegetative bacteria and enveloped viruses. | p21 Household/commercial grade disinfectant wipes Delete |  | Propose moving heirarchy of kill information to be clearly applicable to all disinfectants (from wipes section) |  | Hierarchy of kill information moved | Yes
p21 Residual activity claims |  | Suggest you clarify whether this section applies to hospital grade, commercial/household grade or both. |  |  | Rejected | Yes
p21 Residual activity claims |  |  | Recommend additional information regarding EN 13697 as a preliminary test requirement for PAS 2424. | Agree that EN 13697 is a prerequisite for PAS 2424. This doesn't need to be specified in the document. | Rejected | Yes
p21 Residual activity claims The period over which residual activity against bacteria, yeast or viruses is claimed must not be more than 30 days. For all claims, testing as per those stipulated in methods a) PAS 2424:2014 and b) US EPA Protocol #01-1A and c) ASTM E 1053 (modified) must be performed, with modified abrasion/wear and re-inoculation cycles to represent product usage over the period being claimed, including a final rechallenge step at the end of the period being claimed. |  |  | Possible ambiguity around the 30-day limit, and desire for longer periods. | There does not seem to be particular industry interest in changing this. There are no directly applicable test methods, and currently a 30 day claim requires extrapolation from the existing test methods. Further changes would mean reverting to assessment on a case-by-case basis, which is not an improvement. | Rejected | Yes
Training and Education |  | As included in the TGA Business Plan 2024-25, under strategic objective 2, we believe that Industry would greatly benefit from comprehensive training on the Instructions, to increase awareness about regulatory obligations. As such, we would very much welcome training being provided to Industry, once the amendments have been published (including explanation on microbial hierarchy). | Request training on various regulatory and technical aspects of disinfectants | It is not the TGA's role to provide training on technical matters. The TGA will investigate whether further training is required specifically regarding the Instructions once published. | Information sessions regarding updates to be provided | Yes
Training and Education |  | To further enhance industry’s understanding of the Instructions, it would be valuable for the TGA to run training once the amendments are published. For example, webinars to detail the changes and also the continuing requirements would assist the regulated entities to better understand the requirements and comply. This is also likely to improve the quality of applications lodged with the TGA. Respondent would be happy to promote or support this training. | Request training on various regulatory and technical aspects of disinfectants | It is not the TGA's role to provide training on technical matters. The TGA will investigate whether further training is required specifically regarding the Instructions once published. | Information sessions regarding updates to be provided | Yes
Log reduction and values |  | EPA/TGA methods are qualitative where an overall log reduction is not calculated (59/60 carriers) whereas the EN methodology is quantitative which calculates a log reduction when compared to control. Although a log reduction is not determined for the EPA/TGA methods, control suspension counts and control carrier number are calculated to determine the level of microorganism used to challenge the disinfectant product, these control counts are comparable between EPA/TGA and EN (See below). Due to the additional dilution steps within the EN methods, even if no growth is recorded, a minimum of <0.1 cfu/mL can be achieved therefore a >7 log reduction is the maximum log reduction for EN13697. The acceptance criteria of >4-log reduction in EN13697 supports killing 99.99% of the challenge organism, the levels in efficacy / EN tests are kept quite high when compared to likelihood of seeing these number in domestic surfaces. EN standard states chemical disinfection is the reduction of the number of microorganisms in or on an inanimate or animate matrix, achieved by the irreversible action of a product on their structure or metabolism, to a level judged to be appropriate for a defined purpose, therefore the log reduction stated in the EN standard has been assessed for suitability based on the intended use of the disinfectant. | The log reductions required by the tests are well in excess of real-life situations. | The required log reductions are specified by the test methods. The tests are in-vitro, and may not reflect real-life situations, but are intended to represent worst-case. Note that the methods that are referenced in the document are internationally recognised, and are not under the control of the TGA. | No action required | Yes
Test methods |  | We urge the TGA to consider the acceptability of the Sponsor to be able to use alternative test methods, where studies are scientifically robust, well reported and provided a clear answer as to the efficacy of the product. |  | Other test methods can be used, as long as they are shown to be comparable or better (within the same test method category eg suspension or carrier testing) Sentence added: Although this document provides guidance as to appropriate test methods that may be employed, any test of reasonable scientifically-based, peer reviewed methodology will be considered. | Clarification added. | Yes
Future innovations |  | Technology in this space is always evolving and progressing. Testing requirements should be updated as needed in the Instruction to support innovation. Therefore, we ask that the whole of the Instructions is considered for update with public consultation on a regular basis (e.g. on a five yearly basis). | Request more regular review of the Instructions |  | Acknowledged | Yes
Test methods |  | Urgent and complete review of test methods must be conducted to remove issues such as the sporicidal test methodology. This as an individual example asks (in layman terms) that the product kill more live spores that exist in the test methodology. It is simply impossible to kill over a million spores when the product is only given 100k to start with. Other anomalies exist in other areas as well with a range of pathogens, but most importantly as described above the use of Class 2b as an instrument to unnecessarily stop or restrict products entering or remaining in the market. I also believe that the legislation is being applied in a counter-intuitive way whereby if you exceed the test requirements, then the product is deemed not to have met the requirements of TGO104. This is more an individual assessor issue and I fear may be borne out of personalities and is far too reliant on discretion as opposed to science. | Respondent believes that the referenced test methods do not perform as stated, and may not be conformed with. Not stated, but aware that the individual requests allowing modification of the TGA Disinfectant test method to reduce the contact time | Note that the methods that are referenced in the document are internationally recognised, and are not under the control of the TGA. Respondent has not provided evidence or specifics of concerns around test methods, and it is therefore not possible to investigate further. | Rejected | Yes
Performance of testing |  |  | Respondent suggests that the document contain a requirement for companies to demonstrate that their test method is representative of real-world application of the product. | The test methods referenced are those considered current, and are accepted internationally as suitable for demonstrating effectiveness. | Acknowledged. Rejected. | Yes
Terminology |  | There is an inconsistency in the way log reduction is used in the text it is either log reduction or it is log 10 reduction (sorry can't get the 10 to drop from the line). Consistency is required. Although personally whether the 10 is there or not it is implied. I am looking for consistency to prevent confusion in other readers esp those not in this field. | Respondent suggests consistency is required in using the term log or log10. |  | Edited throughout the document | Yes
Updates to referenced standards |  | I am also assuming that as each standard is released TGA will accept the issue outright as this document does not reference the year of the standards referenced. | Respondent questions whether any edition of the referenced standards may be applied. | (Note that the TGA doesn't check what method version was in effect at the time of testing) Note there is a comment regarding this in Division 2 Section 1 of the document: Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable. | Rejected | Yes