TGA instructions for disinfectant testing

Version 4.0, May 2025

**Copyright**

© Commonwealth of Australia 2024
This work is copyright. You may reproduce the whole or part of this work in unaltered form for your own personal use or, if you are part of an organisation, for internal use within your organisation, but only if you or your organisation do not use the reproduction for any commercial purpose and retain this copyright notice and all disclaimer notices as part of that reproduction. Apart from rights to use as permitted by the *Copyright Act 1968* or allowed by this copyright notice, all other rights are reserved and you are not allowed to reproduce the whole or any part of this work in any way (electronic or otherwise) without first being given specific written permission from the Commonwealth to do so. Requests and inquiries concerning reproduction and rights are to be sent to the TGA Copyright Officer, Therapeutic Goods Administration, PO Box 100, Woden ACT 2606 or emailed to <**tga.copyright@tga.gov.au**>.

Contents

[Introduction 4](#_Toc185596059)

[Part 1 – The TGA disinfectant test 5](#_Toc185596060)

[1. Test 5](#_Toc185596061)

[2. Media 6](#_Toc185596062)

[2.1 Sterile hard water 6](#_Toc185596063)

[2.2 Yeast suspension 6](#_Toc185596064)

[2.3 Medium for growth of test organisms 7](#_Toc185596065)

[2.4 Recovery medium 7](#_Toc185596066)

[3. Test inoculum 7](#_Toc185596067)

[3.1 Test organisms 7](#_Toc185596068)

[3.2 Preparation of inoculum 8](#_Toc185596069)

[3.3 Enumeration of inoculum 8](#_Toc185596070)

[4. Disinfectant dilutions 8](#_Toc185596071)

[5. Temperature 8](#_Toc185596072)

[6. Test procedure 9](#_Toc185596073)

[7. Controls 9](#_Toc185596074)

[7.1 Recovery broth contamination 9](#_Toc185596075)

[7.2 Disinfectant contamination 9](#_Toc185596076)

[7.3 Fertility test 9](#_Toc185596077)

[7.4 Inactivator efficacy 9](#_Toc185596078)

[8. Procedure in case of invalid controls 10](#_Toc185596079)

[9. Results 10](#_Toc185596080)

[Part 2 – Disinfectant performance tests 11](#_Toc185596081)

[Division 1 – Tests for the physical stability, chemical stability and shelf life of disinfectants 11](#_Toc185596082)

[1 Physical stability 11](#_Toc185596083)

[2 Chemical stability 11](#_Toc185596084)

[3 Antimicrobial stability 11](#_Toc185596085)

[4 Stability testing frequency 12](#_Toc185596086)

[5 Packs for stability testing 12](#_Toc185596087)

[6 Prediction of shelf life 12](#_Toc185596088)

[7 Stability program 13](#_Toc185596089)

[Division 2 – Microbial efficacy general requirements 14](#_Toc185596090)

[1. General requirements 14](#_Toc185596091)

[2. Specific tests for hospital grade disinfectants and household / commercial grade disinfectants 15](#_Toc185596092)

[Residual activity claims 21](#_Toc185596093)

[Note: Airborne disinfection technologies 22](#_Toc185596094)

## Introduction

The Australian Government, through the Therapeutic Goods Administration (TGA), is responsible for regulating the quality of therapeutic goods. This is principally achieved by specifying ministerial standards for the goods which may relate to a range of matters (e.g., the quality of the goods and the procedures to be carried out in their manufacture), and otherwise by applying default standards specified in the international pharmacopoeias that are defined in the *Therapeutic Goods Act 1989* (the Act).

[Therapeutic Goods (Standard for Disinfectants and Sanitary Products) (TGO 104) 2019](https://www.legislation.gov.au/F2019L00482/latest/versions) (the Order) (https://www.legislation.gov.au/F2019L00482/latest/versions) is designed to support the quality, safety and efficacy of therapeutic goods that are disinfectants, sanitisers and sanitary fluids and powders.

The Order does so in relation to disinfectants by, in part, specifying a number of important performance requirements. These performance requirements principally require that disinfectants comply with specified microbiological tests, such as the TGA Disinfectant Test or bactericidal standard reference methods as specified in Division 2 to support the claims for bactericidal activity. Additional testing is required where a claim is made for a disinfectant in relation to the product having a sporicidal, fungicidal, tuberculocidal, virucidal or other biocidal use. Other internationally recognised tests, such as those published by ASTM, AOAC or CEN, may be employed to support these claims.

This document is divided into 2 parts that identify and explain the testing requirements for TGO 104:

⦁ The TGA Disinfectant test (Part1)

⦁ Specific testing requirements (Part 2).

A number of terms used in this document, for instance ‘sporicide’, ‘tuberculocide’ are defined in the *Therapeutic Goods Regulations 1990*. A link to these regulations on the Federal Register of Legislation can be found at [Therapeutic Goods Regulations 1990](https://www.legislation.gov.au/F1996B00406/latest/text) (https://www.legislation.gov.au/F1996B00406/latest/text).

The words ‘shall’ and ‘must’ are used interchangeably throughout this document and describe mandatory requirements. The word ‘should’ is used where guidance is provided, but does not preclude other acceptable practices.

It should be noted that hard surface disinfectants and Other Therapeutic Goods (OTG’s) are regulated under Chapter 3 of the Act and are required to meet the applicable standard(s) included therein. Disinfectants and cleaners intended for use on medical devices are regulated as medical devices under Chapter 4 of the Act, and are required to comply with the essential principles identified in Schedule 1 of the [Therapeutic Goods (Medical Devices) Regulations 2002](https://www.legislation.gov.au/F2002B00237/2017-07-01/text) (<https://www.legislation.gov.au/F2002B00237/2017-07-01/text>).

## Part 1 – The TGA disinfectant test

### 1. Test

1. A disinfectant must be tested at the dilution recommended by the manufacturer on the product label.
2. The test consists of the following steps, undertaken in accordance with the instructions or requirements in sections 2 – 8 of the TGA Disinfectant Test (Part 1 of this document) and in the following order:
3. challenging the diluted disinfectant with bacterial inoculum;
4. withdrawing a sample after a given time and culturing the sample in a suitable recovery medium;
5. (except for household/commercial grade disinfectants) after this sampling, challenging the mixture again by a second inoculum, and after a second interval again sampling for culturing.
6. The sample referred to in (2) passes or fails the test according to the criteria specified in section 9 of the TGA Disinfectant Test (Part 1 of this document).

**Table 1. Selection of test parameters for disinfectants using the TGA Disinfectant Test.**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| Class of product | Organisms used in the test | Test option for resuspension of centrifuged organisms | Number of challenges | Inoculum density |
| Hospital grade disinfectant | *Ps. aeruginosa**Pr. Vulgaris**E. coli**S. aureus* | Option A(‘clean’ conditions) | 2 | 2x108 - 2x109 |
| Option B(‘dirty’ conditions) |
| Household/commercial grade disinfectant | *E. coli**S. aureus* | Option C | 1 | 2x108 - 2x109 |

1. The test may be performed:
2. for a hospital grade disinfectant – one of the following:
3. without the addition of sterile yeast as an organic soil (‘Option A’, ‘clean conditions’), using the organisms identified in column 2 of Table 1 and with the number of challenges and inoculum density also identified in Table 1 for such disinfectants; or
4. with the addition of sterile yeast as an organic soil (‘Option B’, ‘dirty conditions’), using the organisms identified in column 2 of Table 1 and with the number of challenged and inoculum density also identified in Table 1 for such disinfectants; or
5. both with and without the addition of sterile yeast as an organic soil, according to the use situations advocated on the label of the product, i.e. when use in both clean (Option A) and dirty (Option B) conditions is claimed.
6. for a household/commercial grade disinfectant, with the use of a nutrient broth as a simulated soil (Option C).

### 2. Media

All media must be contained in capped glass containers and, when stored, the containers must be sealed tightly and refrigerated as necessary.

#### 2.1 Sterile hard water

2.1.1 Dissolve 0.304g anhydrous calcium chloride and 0.065g anhydrous magnesium chloride in distilled water, and make up to one litre.

2.1.2 Dispense into glass containers and sterilize by autoclaving at 121°±1° C for 15 minutes.

**Note:** The AOAC Hard Water formulation[[1]](#footnote-1) may be used.

#### 2.2 Yeast suspension

2.2.1 Weigh 200g of moist compressed baker's yeast. Cream by the gradual addition of sterile hard water using a heavy glass rod for stirring. Decant the creamed portion into a 1L flask, add more water to any lumpy residue remaining and repeat the creaming and decantation until no residue remains and 500mL of water has been used.

2.2.2 Shake the contents of the flask vigorously and strain through a 100-mesh sieve, breaking down any remaining lumps.

2.2.3 Add a further 500mL sterile hard water, shake vigorously and adjust the pH to 6.9-7.1 with 1N Sodium hydroxide.

2.2.4 Transfer 50mL, 100mL or 200mL of the yeast solution into screw capped bottles.

2.2.5 Autoclave at 121°±1°C for 15 minutes and allow the autoclave to cool without releasing pressure. Store cold but not freezing.

2.2.6 Dry two Petri dishes to constant weight. Into each, pipette 25mL of sterilised yeast suspension, and dry to constant weight at 100°C. Calculate the average solids content of the suspension.

2.2.7 Before use, pipette 25mL of the sterilised yeast suspension into a beaker. Determine the pH using a glass electrode, and determine the volume of 1N sodium hydroxide solution needed to adjust the pH to within the range 6.9 to 7.1.

2.2.8 Immediately before use, add to each bottle of sterilised yeast, a volume of sterile hard water and a volume of 1N sodium hydroxide calculated to adjust the concentration of dry yeast to 5.0% and the pH to within the range 6.9 - 7.1 (using the calculation performed in step 2.2.7).

2.2.9 Discard prepared yeast 3 months after preparation.

#### 2.3 Medium for growth of test organisms

2.3.1 Prepare a 10% w/v dextrose solution in distilled water, and sterilise by autoclaving at 121°±1°C for 15 minutes. Cool to room temperature.

2.3.2 Prepare Wright and Mundy medium following the recognised procedure[[2]](#footnote-2) or from a commercial product of the same composition and sterilise by autoclaving at 121°±1°C for 15 minutes. Cool to room temperature.

2.3.3 To each litre of Wright and Mundy medium prepared in item 2.3.2, add 10mL sterile dextrose solution prepared in item 2.3.1.

2.3.4 Aseptically dispense in either 10mL or 15mL amounts, as preferred.

2.3.5 This medium is referred to as ***Wright and Mundy dextrose medium***. AOAC Synthetic Broth[[3]](#footnote-3) may also be used.

#### 2.4 Recovery medium

2.4.1 Prepare nutrient broth as follows or from a commercial product of the same composition:

1. Add the following to 970mL of water and dissolve by heating.

Beef Extract Powder 10g

Peptone 10g

Sodium Chloride 5g

1. Adjust the pH to 8.0-8.4 using 1N Sodium Hydroxide.
2. Boil for 10 minutes and filter. Cool.

2.4.2 To each litre of nutrient broth solution prepared in item 2.4.1 add 30g polysorbate 80.

2.4.3 Adjust pH to 7.2 - 7.4 using 1N Sodium hydroxide.

2.4.4 Autoclave at 121°±1°C for 15 minutes, and immediately shake well to disperse the polysorbate 80.

2.4.5 Dispense aseptically in 10mL amounts into sterile capped glass tubes.

### 3. Test inoculum

#### 3.1 Test organisms

The following 4 organisms, as listed in Table 1, are to be used for hospital grade disinfectants for Option A or Option B.

*Pseudomonas aeruginosa* NCTC 6749;

*Proteus vulgaris* NCTC 4635;

*Escherichia coli* NCTC 8196; and

*Staphylococcus aureus* NCTC 4163.

The following 2 organisms are to be used for household/commercial grade disinfectants for Option C:

*Escherichia coli* NCTC 8196; and

*Staphylococcus aureus* NCTC 4163.

#### 3.2 Preparation of inoculum

3.2.1 Incubate the contents of an ampoule of freeze-dried culture overnight at 37°±1°C in Wright and Mundy dextrose medium.

3.2.2 Inoculate the incubated culture onto nutrient agar (commercially available) slopes in McCartney bottles. Store for up to 3 months at 4°±1°C.

3.2.3 At a suitable period before the test is to be conducted, subculture from an agar slope into 10mL or 15mL quantities of Wright and Mundy dextrose medium. Incubate at 37°±1°C for 24 ± 2 hours.

3.2.4 Subculture from the medium in item 3.2.3 into fresh medium, using an inoculating loop of 4mm in diameter. Incubate at 37°±1°C for 24 ± 2 hours.

3.2.5 Repeat item 3.2.4 daily. For the test procedure use only those cultures which have been subcultured at least 5 times, and not more than 14 times.

3.2.6 Filter test cultures of *P. aeruginosa* and *S. aureus* through sterile Whatmans No. 4 filter papers.

3.2.7 Centrifuge all test cultures until cells are compact, and remove supernatant with a Pasteur pipette.

3.2.8 Resuspend test organisms in the original volume of liquid (i.e. 10mL or 15mL), and shake for 1 minute with a few sterile glass beads.

3.2.8.1 For **Option A**, resuspend in sterile hard water.

3.2.8.2 For **Option B**, resuspend in a mixture of 4 parts yeast suspension (prepared as in item 2.2) to 6 parts sterile hard water.

3.2.8.3 For **Option C**, resuspend in nutrient broth (prepared as in item 2.4 though without the addition of polysorbate 80).

#### 3.3 Enumeration of inoculum

Immediately before testing, sample the resuspended inoculum and enumerate using 10-fold dilutions in quarter-strength Ringer's solution and the pour-plate technique. The number subsequently counted must represent not less than 2 x 108 or more than 2 x 109 organisms per millilitre or the test is considered invalid. Retain tube containing 107 dilution for use in controls (items 7.3 and 7.4).

### 4. Disinfectant dilutions

Quantitatively dilute a sample of the disinfectant to the specified extent, using sterile hard water as diluent. Use not less than 10mL or 10g of sample for the first dilution, and not less than 1mL of any dilution to prepare subsequent dilutions. Make all dilutions in glass containers on the day of testing. The glass containers must be twice rinsed in glass-distilled water, and sterilised.

### 5. Temperature

Where the test environment does not maintain test solutions at 21°±1°C, hold the containers in which the test is to be carried out in a waterbath at this temperature.

### 6. Test procedure

Perform the following test, using each of the test organisms (item 3.1). It is not necessary to test with all organisms simultaneously.

6.1 Add 3mL of diluted disinfectant to a capped glass container.

6.2 Start a timing device. Immediately inoculate disinfectant with 1ml of culture (prepared in item 3.2) and mix by swirling.

6.3 At 8 minutes, subculture one drop (0.02mL ± 0.002mL) into each of 5 tubes containing recovery broth. To ensure delivery of 0.02mL into the first tube of recovery broth at exactly 8 minutes, it will be necessary to withdraw a suitable amount from the disinfectant test mix shortly beforehand. This must be immediately preceded by vortexing. Surplus sample must be returned to the test mix.

6.4 At 10 minutes, inoculate disinfectant with a further 1ml of culture, and mix by vortexing.

6.5 At 18 minutes, repeat the process as described in item 6.3.

6.6 Mix the contents of all tubes of recovery broth by vortexing. Incubate at 37°±1°C for 48 ± 2 hours.

6.7 Examine for growth and record results.

6.8 For each test organism repeat steps 6.1 - 6.7 on each of 2 subsequent days, using a fresh disinfectant dilution and a freshly prepared bacterial suspension.

### 7. Controls

#### 7.1 Recovery broth contamination

Incubate one uninoculated tube of recovery broth at 37°±1°C for 48 ± 2 hours and examine for growth. If growth occurs, the test is considered invalid due to contamination of the recovery broth.

#### 7.2 Disinfectant contamination

To 1 tube of recovery broth, add 0.02mL of diluted disinfectant. Incubate at 37°±1°C for 48 ± 2 hours. If growth occurs, the test is considered invalid. Growth in item 7.2 but not item 7.1 indicates contamination of the disinfectant test solution.

#### 7.3 Fertility test

To 1 tube of recovery broth, add 1.0mL of the 107 dilution of the test organism, retained in item 3.3. Incubate at 37°±1°C for 48 ± 2 hours and examine for growth. If no growth occurs, the test is considered invalid.

#### 7.4 Inactivator efficacy

To 1 tube of recovery broth, add 0.02mL of diluted disinfectant and 1.0mL of the 107 dilution retained in item 3.3. Incubate at 37°±1°C for 48 ± 2 hours, and examine for growth. If no growth occurs, the test is considered invalid. Growth in item 7.3 but not in item 7.4 indicates inadequate inactivation of the disinfectant.

### 8. Procedure in case of invalid controls

When any control renders the test invalid, the test is to be repeated. Fresh recovery broth is to be used if growth occurred in control item 7.1, or if no growth occurred in control items 7.3 or 7.4.

Should disinfectant contamination be indicated by control item 7.2 on both occasions, the disinfectant is considered to fail the test. Should inadequate inactivation of the disinfectant be indicated by control item 7.4 on both occasions, the test is considered invalid.

### 9. Results

The dilution test is considered to pass the test if there is no apparent growth in at least two out of the five recovery broths specified in item 6.3 and no apparent growth in at least two of the five recovery broths specified in item 6.5 on all three occasions, using all test organisms.

## Part 2 – Disinfectant performance tests

### Division 1 – Tests for the physical stability, chemical stability and shelf life of disinfectants

#### 1 Physical stability

Physical stability testing may include:

1. appearance (e.g. emulsion stability, clarity etc.)
2. odour
3. pH
4. immediate container and product compatibility

#### 2 Chemical stability

Chemical stability testing must comply with the following:

1. active(s), to be determined by an assay that is valid for the active (eg. HPLC, GC, titration),
2. active(s) to be above the minimum level required to pass the appropriate efficacy test(s) of the product class as described in Division 2 of this document over the shelf life, and
3. all chemical testing to be carried out on duplicate samples.

#### 3 Antimicrobial stability

A disinfectant must pass at least one suitably sensitive test according to the appropriate evaluation guidelines for the grade of disinfectant (see Table 2 below) at the end point at which testing is occurring, with the levels of active(s) at the predicted final level.

**Table 2. Appropriate tests according to disinfectant grade.**

|  |  |
| --- | --- |
| Disinfectant | Appropriate test |
| Hospital Grade | Carrier test[[4]](#footnote-4) (with or without soil as appropriate) |
| Household/Commercial Grade | The TGA disinfectant Test in Part 1, under the conditions in Option C, EN 1276[[5]](#footnote-5), or carrier test (with organic soil)[[6]](#footnote-6) |

#### 4 Stability testing frequency

##### Accelerated studies

Suggested testing intervals are 0, 3, 6 and 12 months, but other time points e.g. 1, 2, or 5 months, may also be used and may be necessary for adequate accelerated studies. The minimum active concentration level from which the shelf life is predicted from accelerated data must be no lower than the minimum active concentration required to pass the appropriate microbial test.

##### Real time studies

Testing at initial and annual intervals should be sufficient. However, it is not necessary to have completed these testing programs before launching the product. Samples from production should undergo real time testing. These samples should be monitored over the shelf life of the product as proposed by the sponsor.

#### 5 Packs for stability testing

Stability testing should be carried out on product stored in the proposed packaging material, and in the case of real time testing, this requirement is mandatory.

#### 6 Prediction of shelf life

Shelf life may be predicted from accelerated data only if the accuracy and reproducibility of the results is adequate to support the extrapolation. All data must be fully defensible.

Extrapolation at various times and temperatures may be determined according to the general rules described in Table 3 below.

 **Table 3. Shelf life extrapolation.**

|  |  |  |
| --- | --- | --- |
| Elevated Temp(a) above storage conditions(b) | Time Period | Possible Shelf Life Prediction |
| +10°C | 3 months | 1 year |
| +15°C | 3 months | 18 months |
| +10°C | 6 months | 2 years |
| +15°C | 6 months | 3 years |
| +10°C | 9 months | 3 years |

(a) incubator temperatures must be monitored and logged

(b) storage conditions for Australia are considered to be the temperatures of 25 - 30°C. The reference temperature used for stability testing will be that on the label.

Alternative predictive models/rules for extrapolation can be used, however, after proper scientific validation.

For all predicted shelf life studies, at least four real time data points, including the initial and three months, should be evaluated using acceptable statistical methods to justify the extrapolation.

The predicted value of the final concentration of the disinfectant active should not fall below the 90% limit of the label claim at the end of the extrapolated shelf life.

#### 7 Stability program

The stability program should collect real time data from a production batch stored at the recommended storage condition. This should be used to confirm the extrapolated shelf life.

The TGA should be advised if the product falls outside specification, with appropriate action, commensurate with public safety, negotiated and agreed.

If there are minor changes to excipients, product stability should be confirmed, including a single microbial efficacy test as for the final determination of shelf life, unless justification is provided using a risk-based approach.

### Division 2 – Microbial efficacy general requirements

#### 1. General requirements

All tests should be carried out by a GMP licensed laboratory or laboratory accredited to ISO/IEC 17025 or equivalent. Data sourced from non-accredited laboratories will be accepted in relation to testing conducted using blood borne viruses, including HIV, Hepatitis B and C, Ebola and other haemorrhagic viruses, but not generally in other instances.

Although this document provides guidance as to appropriate test methods that may be employed, any test of reasonable scientifically-based, peer reviewed methodology will be considered.

Where international standards and test methods (CEN, AOAC, US EPA, ASTM) are referenced in testing, testing should be conducted against the current and relevant standard at the time the testing was conducted, and the edition number of the standard clearly stated within the testing results. Where testing has been conducted to a more recent standard than those set out below, the more recent standard will be acceptable. Testing should be performed according to the test method used with regard to test organisms, number of carriers tested, and type and concentration of organic soil to be used to simulate dirty and clean conditions. Exceptions or additional requirements are specified in this document. Testing should also reflect the in-use conditions of the product, within the limits of the test method.

Some of the test methods referenced in this document specify that testing should be performed at three product concentrations. This is applicable to the product development phase, and testing at the in-use concentration of the product only is expected once a product is out of development.

Regardless of whether the product is used undiluted or diluted by the user to the Minimum Recommended Concentration, the product should be formulated at the lower extreme of the active ingredient in the product specification, stored to expiration and then subjected to the microbial efficacy tests, and this data should be available on request.

The product may be tested before the end of the shelf life to initially provide a full set of data for evaluation purposes if product stored to the end of the expiration period is not available. The data can be generated from a single batch of product. However, data from further batches may be requested by the TGA in certain circumstances (i.e., the product contains a novel active or the single batch stability data raised concerns). Testing at the end of the shelf life will then need to be performed according to the microbiological stability requirements specified in Section 4 above, once product stored to the end of the expiration period becomes available. End of shelf-life data should be generated from a minimum of two batches.

A common approach to developing a worst-case product for testing is to use accelerated life testing to predict the final level of actives. These levels (with a safety factor) are then used in the formulation of a sample for efficacy testing. However, some products are unstable at elevated temperatures and may not be suitable for accelerated stability testing. In general, real time stability studies are preferred.

Data from all tests specified for each grade and level of disinfectant is not required for monitoring of stability (see Table 3 above for the appropriate tests to be used to generate stability data). Tests must be carried out at the pH, temperature and time (unless the stability studies are conducted under the accelerated conditions) recommended on the label for the product use, or each level of use where there is more than one level.

Test methods must be validated by individual laboratories for each test method used in accordance with tests which have been validated or refereed at national or international level. This should involve validation of individual operators by performance of the specified test using a control or product that is known to pass the specified test at a given concentration as well as testing to determine the most appropriate neutraliser/inactivator, if necessary.

The most stringent claim is regarded as the organism least susceptible to the test product. In general, the hierarchy of susceptibility from hardest to easiest to kill is: bacterial spores, mycobacteria, non-enveloped viruses, fungi, vegetative bacteria and enveloped viruses.

#### 2. Specific tests for hospital grade disinfectants and household / commercial grade disinfectants

##### Hospital grade disinfectants

Bactericidal efficacy (excluding tuberculocidal) is the only mandatory requirement for a hospital grade disinfectant without specific claims.

All testing should be performed using the in-use product conditions.

###### Bactericidal testing

These disinfectants MUST pass Option B of the TGA Disinfectant Test under dirty conditions (as per Table 1 in Part 1 of this document). This is a semi-quantitative suspension test. Testing in accordance with the dirty conditions option of EN 13727[[7]](#footnote-7) will also be accepted. Other suspension tests may be acceptable if shown to be equivalent to the above methods. AOAC Hard water or equivalent should be used as necessary.

If the disinfectant is clearly labelled for use on a pre-cleaned surface, Option A of the TGA Disinfectant Test may be used. Under these circumstances, the clean conditions option of EN 13727 will also be accepted as an evaluation test. Other suspension tests may be acceptable under clean conditions provided AOAC Hard Water or equivalent is used.

If specific label claims are made for vegetative bacterial species other than those covered by the required test organisms, the testing above must be conducted using these additional organisms. Specific testing as described below must be conducted when label claims are made for activity against mycobacteria and spore-forming organisms.

These disinfectants MUST also pass a bactericidal carrier test. AOAC[[8]](#footnote-8),[[9]](#footnote-9) methodology or a quantitative carrier test (methods such as ASTM E2197[[10]](#footnote-10), ASTM E2111[[11]](#footnote-11), EN 14561[[12]](#footnote-12), EN 13697[[13]](#footnote-13) or EN 17387[[14]](#footnote-14)) may be used.

|  |  |
| --- | --- |
| Information | **NOTE:** EN 13697 is specified as a method for the evaluation of disinfectants for use in food, industrial, domestic and institutional areas. It is not applicable for the evaluation of disinfectants for medical or clinical applications. However, as EN 17387 was only released in 2021, EN 13697 will continue to be accepted as an appropriate test method at this point. |

If the AOAC Use Dilution Test is chosen, the disinfectant passes when there is no growth in 59 out of 60 carriers per organism. If additional claims against organisms not included in the original test are to be made (e.g. ‘Kills E coli 0157’), the extra organism(s) can be tested using 10 carriers rather than 60 when using the AOAC test methods. None of the 10 carriers should show growth. Soil should be included at a minimum of 5% blood serum and inorganic soil such as AOAC hard water.

If the Hard Surface Carrier Test is used, the criteria for pass or fail are described in the test method. If the ASTM or EN test methods are used, applicable test conditions are as described in the test method, however a minimum 5-log10 reduction must be achieved.

If any specific biocidal claims are made (i.e., virucidal, fungicidal, tuberculocidal, sporicidal, or other biocidal activity), the disinfectant must pass appropriate tests as specified below. All tests should be carried out using the exposure time, temperature and pH specified on the label. For products intended for use on surfaces that have not been pre-cleaned, 5% organic soil must be included. For products that are intended for use on surfaces that have been pre-cleaned, organic soil need not be added to the test, with the exception of testing against blood-borne viruses, as discussed in the virucidal claim section below.

###### Fungicidal testing

For a fungicidal claim, any test of reasonable scientifically-based, peer reviewed methodology will be considered, although carrier test methodology is preferred. Acceptable tests include the AOAC Use Dilution Test, which can be modified for fungi, with 10 carriers for each of two batches of product, and acceptance criteria of all fungal spores on all carriers killed, the AOAC Fungicidal Test[[15]](#footnote-15), which can be modified for a carrier test, EN 17387, and the ASTM E2197 test methodology, with acceptance criteria as for the modified AOAC Use Dilution Test. Tests conducted in accordance with EN 13624 or EN 13697 will also be accepted.

###### Virucidal testing

For a general virucidal claim, not including blood borne viruses such as HIV, Hepatitis B (HBV), Hepatitis C (HCV) or Ebola, the disinfectant MUST pass tests using either Parvovirus or a Poliovirus surrogate as the test viruses. Possible alternatives for Poliovirus include Hepatitis A, Feline Calicivirus or Murine Norovirus. Testing to support label claims against HIV, Hepatitis B (HBV), or Hepatitis C (HCV) must be conducted with the organic soil component consisting of at least 50% whole blood.

The tests may be suspension tests but carrier tests are preferred. Methods that may be used as a basis are the AOAC Use Dilution Test modified for viruses, EN 16777[[16]](#footnote-16) and ASTM E2197. Guidance on carrier test methodology is provided in ASTM E 1053[[17]](#footnote-17). Guidance on suspension test methodology is provided in ASTM E 1052[[18]](#footnote-18). If a suspension test is used, the methodology of EN 14476[[19]](#footnote-19) is acceptable, if the test is conducted under dirty conditions.

Viral recovery systems that may be used include tissue culture, embryonated egg and animal inoculation.

Tests on the designated prototype viruses should be performed in quadruplicate against a recoverable viral titre of at least 4-log10, which must be recoverable from the test surface or suspension, and should show complete viral inactivation. If cytotoxicity is apparent, a 3-log10 reduction must be demonstrated beyond the cytotoxic level and there should be complete viral inactivation. Cytotoxicity is more easily overcome in suspension tests, for which there should be a 4-log10 reduction with complete viral inactivation. Suitable controls should be employed, which include:

* cytotoxicity controls
* disinfectant neutralization controls
* quantitative viability control
* cell control
* carrier wash-off control (to determine the efficiency of removal of the test organism from the carrier).

If a label claim against HIV, HBV, HCV or other specific virus is made, separate data must be provided, in addition to the above. Suspension tests may be used for additional specific viral claims, but the recoverable viral challenge should reflect the titre found in the clinical situation. Complete viral inactivation is preferred, with a minimum 4-log10 reduction of specific viruses. For HIV, the method should be based on the principles described above. The use of a surrogate virus for HIV claims is not permitted.

For viruses that cannot be cultured, surrogate viruses may be used. For HBV, tests with a surrogate organism such as Duck Hepatitis B Virus (DHBV) are acceptable. The electron microscope viral disruption and antigenicity tests are NOT acceptable, as they are not reliable indicators of HBV infectivity. A suitable surrogate for HCV may be Bovine Viral Diarrhoeal Virus (BVDV). Tests for HIV and HCV MUST be conducted against cell-associated virus – a suspension test may be used if necessary. The organic soil component for tests against HIV and surrogates for HBV and HCV and Ebola MUST be a minimum of 50% whole blood. Suitable references for test methods are Murray (1991)[[20]](#footnote-20), Druce (1995)[[21]](#footnote-21), and Lavelle (1987)[[22]](#footnote-22).

With regard to testing against HIV, there can be difficulties associated with testing using cell-associated virus. Non-cell-associated virus can be used for HIV testing of hospital or household grade disinfectants not intended for use on medical devices if:

* a carrier test is used instead of a suspension test,
* the minimum level of 50% of the organic soil used in the test consists of whole blood, and
* the product is clearly labelled for use on a pre-cleaned surface.

There are a number of other viruses which cannot be cultured which may be the subject of label claims. These include Norwalk virus and Norovirus. There are also viruses such as SARS-associated coronavirus, which can be cultured but may not be suitable for use in test laboratories due to biosecurity issues. Feline calicivirus has been accepted as a surrogate to justify claims against Norwalk and Noroviruses. For claims against SARS or SARS-CoV-2 (COVID-19), Human coronavirus 229E or Murine hepatitis virus can be used as a surrogate if either the SARS virus or the COVID-19 virus cannot be used.

Other strains of Human coronavirus will also be considered if Human coronavirus 229E cannot be used. In the event that Human coronavirus or Murine hepatitis virus cannot be used, consideration will be given to use of other animal coronaviruses. Viruses that have been suggested include Bovine coronavirus and Feline coronavirus.

If virucidal testing is limited to lipid/enveloped viruses, such as Herpes simplex virus, a label claim for general virucidal activity will not be permitted. The label must reflect the specific viruses used for the limited testing.

###### Tuberculocidal/mycobactericidal testing

For a tuberculocidal claim, results from a quantitative carrier test should be provided. A variety of organisms may be used, such as:

*M bovis* *(BCG)*

*M tuberculosis* *H37RV*

*M terrae* *ATCC 15755*

|  |  |
| --- | --- |
| Information | **NOTE:** *M smegmatis* is NOT acceptable as this organism is comparatively easy to kill and its resistance patterns are substantially different from the organisms listed above. |

The acceptance criterion is a 6-log10 reduction in test organisms. If the carrier test used is not quantitative (e.g. the AOAC *Test for Tuberculocidal Activity of Disinfectants[[23]](#footnote-23))*, a suspension test may be used to demonstrate a 6-log10 reduction.

Appropriate tuberculocidal tests that may be modified in line with points raised above include:

* the AOAC *Test for Tuberculocidal Activity of Disinfectants* (with the exception of the in-vitro screening test using *M smegmatis*)
* the EPA Quantitative Tuberculocidal Activity Test[[24]](#footnote-24)
* ASTM E2111
* ASTM E2197
* the Ascenzi test[[25]](#footnote-25)
* EN 14348[[26]](#footnote-26).

A qualitative or semi-quantitative test method may be used, if a 6-log10 reduction in mycobacteria can be extrapolated from the results.

###### Sporicidal testing

For a sporicidal claim, a quantitative carrier test or a suspension test should be used. The results should show a 6-log10 reduction in spores. Carrier tests that may be suitable are those performed by the Hospital Infection Research Laboratory at Dudley Road Hospital, Birmingham, UK[[27]](#footnote-27), and tests conducted in accordance with ASTM E2197. The methodology of EN 17126[[28]](#footnote-28) will be accepted, if the test is modified to show a 6-log10 reduction in spores over the labelled exposure time. A qualitative or semi-quantitative test method (such as the AOAC Sporicidal Test[[29]](#footnote-29)) may be used, if a 6-log10 reduction in spores can be extrapolated from the results, and only in the absence of organism recovery/growth.

Suitable test organisms include:

*Clostridium sporogenes* ATCC 3584

*Bacillus subtilis* ATCC 19659 or NCTC 10073

Label claims against *Clostridium difficile* can only be made if the methodology uses spores. The claim cannot be made if a vegetative form of *C. difficile* is used.

Requirements specific to particular presentations of hospital-grade disinfectants can be found below. These requirements are in addition to those above.

###### Hospital grade disinfectant wipes

The following tests are intended to apply to products making claims of surface disinfection. They are not intended to apply to products claiming activity solely within the wipe.

A hospital grade disinfectant wipe is required to:

1. Pass a bactericidal suspension test such as the TGA Disinfectant Test (Option B for dirty conditions, Option A for clean conditions), or EN 13727. The suspension test chosen should be conducted on the product after extraction from the wipe.
2. Alternatively, a modified AOAC Germicidal Spray Test may be performed with 60 carriers (or 10 carriers for supplemental organisms), tested by wiping the surface of the carriers with the saturated wipe and subculturing the carriers after the specified contact time. One wipe should be used for a minimum of 10 carriers. The performance criteria should be no growth from 59/60 carriers. Testing in accordance with EN 16615[[30]](#footnote-30) will also be accepted.
3. If specific biocidal claims are made, these claims should be supported with data from a suspension or carrier test, with added soil as appropriate. If a suspension test is chosen, the test should be conducted on the product after extraction from the wipes. These tests should be carried out according to the general requirements of hospital grade disinfectants above.
4. In addition to verifying all specific biocidal claims with a carrier or suspension test, the product must also pass a simulated in-use test (as described above) against the most stringent claim being made, to demonstrate that the efficacy of the disinfectant is not reduced when combined with a wipe.
5. Where a wipe is reusable, the wipe must pass a simulated in-use test after the product has been subjected to the re-use protocol. Any reasonably designed test will be considered – a suggested method would include challenging the wipe by wiping it on a variety of surfaces. These surfaces should have been subject to periodic contamination with a microbiological bioburden at a concentration of 106 CFU/5mL use solution. The wipe should be allowed to dry between uses. At the end of the use period, the product should be subjected to a simulated in-use test as described in clause 4 above.

###### Hospital grade disinfectant wipes and sponges where the ingredient that is active cannot be expressed

Any reasonably designed test will be considered – a suggested method would include challenging the wipe or sponge by wiping it on a variety of surfaces. These surfaces should have been subject to periodic contamination with a microbiological bioburden. The wipe or sponge should be allowed to dry between uses. At the end of the use period, the product should be subjected to a simulated in-use test as described in clause 4 of ‘Hospital grade disinfectant wipes’ above. If claims against specific bacteria are made, then these organisms should be included in the test. Testing in accordance with EN 16615 will also be accepted.

###### Hospital grade disinfectant surface sprays

A hospital grade disinfectant that is a surface spray disinfectant is required to:

1. Pass a bactericidal carrier test such as the AOAC Germicidal Spray Test. Sixty carriers per organism should be used, plus 10 carriers for each additional bactericidal claim. The performance criteria should be no growth from 59/60 carriers and 10/10 of the carriers for the additional claims. Testing in accordance with EN 14561, EN 13697 or EN 17387 will be accepted (the test to be chosen depends on the intended use – for hospital grade disinfectants, EN 14561 is the most appropriate).
2. Pass the tests specified in Clause 1 with added organic and inorganic soil (if the product is for dilution). For the AOAC test, organic soil should be a minimum of 5% blood serum and inorganic soil should be hard water (AOAC Hard Water or equivalent) For products intended for use on pre-cleaned surfaces, organic soil need not be included. For EN tests, organic and inorganic soil is added as described in the method. A surface spray disinfectant should be tested against *Salmonella enterica* serotype *choleraesuis* (*Salmonella choleraesuis*), *Staphylococcus aureus* and *Pseudomonas aeruginosa* if the AOAC Germicidal Spray Test is used. If other tests are chosen, it is expected that organisms similar to these would be used.
3. Pass carrier tests as described in previous clauses, if specific biocidal activities are claimed. Inorganic and organic soil should be included, unless the product is clearly labelled for use on pre-cleaned surfaces.

##### Household/commercial grade disinfectants

Household/commercial grade disinfectants must:

1. Pass Option C of the TGA Test, or a test conducted in accordance with EN 1276, or
2. Pass an appropriate bactericidal carrier test, using AOAC methodology (60 carriers per stated organism and 10 carriers for each additional organism.), EN 13727 or EN 13697. Organic soil is not necessary but inorganic soil such as AOAC Hard Water or equivalent should be included.

A household/commercial grade disinfectant must also pass tests as for hospital grade disinfectants where specific biocidal activities are claimed. Inorganic soil should be included. Note: If sporicidal claims are made, testing in accordance with EN 13704[[31]](#footnote-31) will be accepted.

Requirements specific to particular presentations of household/commercial grade disinfectants can be found below. These requirements are in addition to those above.

###### Household/commercial grade disinfectant wipes

The following tests are intended to apply to products making claims of surface disinfection. They are not intended to apply to products claiming activity solely within the wipe.

A household/commercial grade disinfectant wipe is required to:

1. Pass a bactericidal suspension test such as the TGA Test Option C or testing in accordance with EN 1276. The suspension test chosen should be conducted on the product after extraction from the wipe. Organic soil is not necessary but inorganic soil such as hard water should be included.
2. Alternatively, a modified AOAC Germicidal Spray Test as described for hospital grade disinfectants may be used. Testing in accordance with EN 16615 will also be accepted.
3. If specific biocidal claims are made, all claims should be supported with data from a suspension or carrier test. If a suspension test is chosen, the test should be conducted on the product after extraction from the wipes. These tests should be carried out according to the requirements of household/commercial grade disinfectants as above.
4. In addition to verifying all biocidal claims with a carrier or suspension test, the product must also pass a simulated in-use test (as described above) against the most stringent claim being made, to demonstrate that the efficacy of the disinfectant is not reduced when combined with a wipe.
5. Where a wipe is reusable, the wipe must pass a simulated in-use test after the product has been subjected to a re-use protocol. Any reasonably designed test will be considered – a suggested method would include challenging the wipe by wiping it on a variety of surfaces. These surfaces should have been subject to periodic contamination with a microbiological bioburden at a concentration of 106 CFU/5mL use solution. The wipe should be allowed to dry between uses. At the end of the use period, the product should be subjected to a simulated in-use test as described in clause 4 above.

###### Household/commercial grade disinfectant wipes and sponges where the ingredient that is active cannot be expressed

Any reasonably designed test will be considered – a suggested method would include challenging the wipe or sponge by wiping it on a variety of surfaces. These surfaces should have been subject to periodic contamination with a microbiological bioburden. The wipe or sponge should be allowed to dry between uses. At the end of the use period, the product should be subjected to a simulated in-use test as described above. If claims against specific bacteria are made, then these organisms should be included in the test. Testing in accordance with EN 16615 will also be accepted.

###### Household/commercial grade disinfectants that are surface spray disinfectants

A household/commercial grade disinfectant that is a surface spray disinfectant is required to:

1. Pass a bactericidal carrier test such as the AOAC Germicidal Spray Test. Sixty carriers per organism should be used, plus 10 carriers for each additional bactericidal claim. The performance criteria should be no growth from 59/60 carriers. Testing in accordance with EN 14561, EN 13697 or EN 17387 will be accepted (test to be chosen depends on intended use).
2. Pass the tests specified in clause 1 with added inorganic soil, if intended as a surface spray disinfectant that is diluted before use. A surface spray disinfectant should be tested against *Salmonella enterica* serotype *choleraesuis* (*Salmonella choleraesuis*) *and Staphylococcus aureus* if the AOAC Germicidal Spray Test is used. If other tests are chosen, it is expected that organisms similar to these would be used.
3. Pass carrier tests as described in previous clauses, if specific biocidal activities are claimed. Inorganic soil should be included.

#### Residual activity claims

1. If any claim of residual activity is made, the residual activity must be established using the methodology in:

(a) for bacteria or yeast—PAS 2424:2014 *Quantitative surface test for the evaluation of residual antimicrobial (bactericidal and/or yeasticidal) efficacy of liquid chemical disinfectants on hard non-porous surfaces. Test method*; or

(b) for bacteria—US EPA Protocol #01-1A *Protocol for Residual Self-Sanitising Activity of Dried Chemical Residues on Hard, Non-Porous Surfaces*; or

(c) for viruses— ASTM E 1053 relating to a virucidal carrier test modified to include the abrasion and re-inoculation methodology in (a), or the wear and re-inoculation methodology in (b) would be considered. There is currently no standard method for a residual activity claim for a virus. Suitable virucidal efficacy test controls should be employed, including cytotoxicity controls, disinfectant neutralisation controls, quantitative viability control, cell control and carrier wash-off control. Guidance on carrier test methodology is provided in ASTM E 1053.

(2) A claim of residual activity must be substantiated by test data that:

(a) meets acceptance criteria set at a 3-log10 difference between the test product and control product in their capability to continue to produce a reduction in the number of viable cells of relevant test organisms; and

(b) is derived using test methods justified with regards to the intended use of the disinfectant.

(3) The period over which residual activity against bacteria, yeast or viruses is claimed must not be more than 30 days. For all claims, testing as per those stipulated in methods a) PAS 242:2014 and b) US EPA Protocol #01-1A and c) ASTM E 1053 (modified) must be performed, with modified abrasion/wear and re-inoculation cycles to represent product usage over the period being claimed, including a final rechallenge step at the end of the period being claimed.

#### Note: Airborne disinfection technologies

Airborne disinfection, often automated, is a widely used disinfection method in hospitals and healthcare settings throughout the world, including Australia. Examples of such technologies include hydrogen peroxide vapour, or hydrogen peroxide + peracetic acid fogging. Testing as described above may not be applicable to these types of technologies.

European standard EN 17272 - *Quantitative Carrier test for Airborne Room Disinfection by Automated Processes - Determination of Bactericidal, Fungicidal, Yeasticidal, Sporicidal, Tuberculocidal, Mycobactericidal, Virucidal and Phagocidal Activities in the Medical Area, Veterinary Area and Food, Industrial, Domestic and Institutional Areas - Test Methods and Requirements (Phase 2, Step 2)* is an appropriate test method. In addition, the French standard NFT 72 281:2014 *Methods of airborne disinfection of surfaces - determination of bactericidal, fungicidal, yeasticidal, mycobactericidal, tuberculocidal sporicidal and virucidal activity, including bacteriophages* is an acceptable test method.

Automated airborne disinfection systems also require in-use validation using biological or chemical indicators, with 6-log10 *Geobacillus stearothermophilus* globally accepted as the microorganism challenge to be used. The US EPA Protocol for Room Sterilization by Fogger Application (<https://www.epa.gov/sites/production/files/2015-09/documents/room-sterilization.pdf>) is an acceptable method for this testing.

Version history

|  |  |  |  |
| --- | --- | --- | --- |
| Version | Description of change | Author | Effective date |
| V4.0 | Full technical review | Microbiology Section, HPRG | May 2025 |

1. US EPA Standard Operating Procedure for Preparation of AOAC and OECD hard water and other diluents for preparation of antimicrobial products SOP Number: MB-30-01 Date Revised: 02-03-17 <https://www.epa.gov/sites/production/files/2018-01/documents/mb-30-01.pdf> [↑](#footnote-ref-1)
2. Refer to *Handbook of Microbiological Media* by Ronald M. Atlas and edited by Lawrence C. Parks, CRC Press, 1993, where it is titled *Synthetic Broth, AOAC.* [↑](#footnote-ref-2)
3. Synthetic Broth, AOAC is a chemically defined medium recommended by AOAC for growing inoculum, making subcultures and preparing various dilutions while testing disinfectants. [↑](#footnote-ref-3)
4. Disinfectants are used on contaminated surfaces, thus making it necessary to evaluate their microbicidal action on representative carrier materials contaminated with a dried challenge microorganism(s). These tests are more stringent than suspension tests, such as the TGA Disinfectant Test. Information on appropriate carrier tests is provided in Division 2, section 2. [↑](#footnote-ref-4)
5. European Standard EN 1276 *Chemical disinfectants and antiseptics - Quantitative suspension test for the evaluation of bactericidal activity of chemical disinfectants and antiseptics used in food, industrial, domestic and institutional areas - Test method and requirements (phase 2, step 1)* [↑](#footnote-ref-5)
6. Chemical stability alone is not sufficient evidence of product stability. Efficacy testing is required to be included in stability studies to account for such effects as interactions with the container, effects of storage, and transport conditions. Chemical stability also does not account for efficacy under usage conditions, such as the use of the product on surfaces, or use at the intended dilution. [↑](#footnote-ref-6)
7. European Standard EN 13727 *Chemical disinfectants and antiseptics – Quantitative suspension test for the evaluation of bactericidal activity in the medical area – Test method and requirements (phase 2, step 1)* [↑](#footnote-ref-7)
8. *Official Methods of Analysis* *of the AOAC International*, current edition, Use-Dilution Method 955.15 and 964.02 [↑](#footnote-ref-8)
9. *Official Methods of Analysis of the AOAC International*, current edition, Hard Surface Carrier Test Methods [↑](#footnote-ref-9)
10. ASTM E2197 *Standard Quantitative Disk Carrier Test Method for Determining the Bactericidal, Virucidal, Fungicidal, Mycobactericidal and Sporicidal Activities of Liquid Chemical Germicides* [↑](#footnote-ref-10)
11. ASTM E2111 *Standard Quantitative Carrier Test Method to Evaluate the Bactericidal, Fungicidal, Mycobactericidal and Sporicidal Potencies of Liquid Chemicals* [↑](#footnote-ref-11)
12. European Standard EN 14561 *Chemical disinfectants and antiseptics - Quantitative carrier test for the evaluation of bactericidal activity for instruments used in the medical area - Test method and requirements (phase 2, step 2)* [↑](#footnote-ref-12)
13. European Standard EN 13697 *Chemical disinfectants and antiseptics - Quantitative non-porous surface test*

*for the evaluation of bactericidal and/or fungicidal activity of chemical disinfectants used in food, industrial,*

*domestic and institutional areas - Test method and requirements without mechanical action (phase 2, step 2)* [↑](#footnote-ref-13)
14. European Standard EN 17387 *Chemical disinfectants and antiseptics - Quantitative test for the evaluation of bactericidal and yeasticidal and/or fungicidal activity of chemical disinfectants in the medical area on nonporous surfaces without mechanical action - Test method and requirements (phase 2, step 2)* [↑](#footnote-ref-14)
15. *Official Methods of Analysis of the AOAC International*, current edition, Fungicidal Activity of Disinfectants Method 955.17 [↑](#footnote-ref-15)
16. European Standard EN 16777 *Chemical disinfectants and antiseptics. Quantitative non-porous surface test without mechanical action for the evaluation of virucidal activity of chemical disinfectants used in the medical area. Test method and requirements (phase 2/step 2)* [↑](#footnote-ref-16)
17. ASTM E 1053*Standard Test Method to Assess Virucidal Activity of Chemicals Intended for Disinfection of Inanimate, Nonporous Environmental Surfaces* [↑](#footnote-ref-17)
18. ASTM E 1052 *Standard Test Method to Assess the Activity of Microbicides against Viruses in Suspension* [↑](#footnote-ref-18)
19. European Standard EN 14476 *Chemical disinfectants and antiseptics – Quantitative suspension test for the evaluation of virucidal activity in the medical area – Test method and requirements (phase 2, step 1)* [↑](#footnote-ref-19)
20. Murray SM, JS Freiman, K. Vickery, D.Lim, YE Cossart, RK Whiteley ‘Duck Hepatitis B Virus: a model to assess efficacy of disinfectants against hepadnavirus activity’. *Epidemiol Infect* (1991) 106, 435 – 443 [↑](#footnote-ref-20)
21. Druce JD, D Jardine, SA Locarnini, CJ Birch ‘Susceptibility of HIV to inactivation by disinfectants and ultraviolet light’ *Journal of Hospital Infection* (1995) 30 pp [↑](#footnote-ref-21)
22. Lavelle George C, ‘Virucidal activity of Disinfectants: Predicting and Assessing Product Efficacy’ *Chemical Times and Trends*, January 1987 45-50 [↑](#footnote-ref-22)
23. *Official Methods of Analysis of the AOAC International*, current edition, Tuberculocidal Activity of Disinfectants Method 965.12 [↑](#footnote-ref-23)
24. US EPA *Standard Operating Procedure for Quantitative Suspension Test Method for Determining Tuberculocidal Efficacy of Disinfectants Against Mycobacterium bovis (BCG)*, SOP Number: MB-16-01, 2009 [↑](#footnote-ref-24)
25. Ascenzi JM, RJ Ezzell, TM Wendt, ‘A More Accurate Method for Measurement of Tuberculocidal Activity of Disinfectants’ *Applied and Environmental Microbiology*, Sept 1987 p2189-2192 [↑](#footnote-ref-25)
26. European Standard EN 14348 *Chemical disinfectants and Antiseptics – Quantitative Suspension Test for the Evaluation of Mycobactericidal Activity of Chemical Disinfectants in the Medical Area including Instrument Disinfectants – test methods and requirements (phase 2, step 1)* [↑](#footnote-ref-26)
27. ‘Sporicidal Activity of Glutaraldehyde and Hypochlorites’ *Journal of Hospital Infection*, (1980) 1 63-75 [↑](#footnote-ref-27)
28. European Standard 17126 Chemical disinfectants and antiseptics. Quantitative suspension test for the evaluation of sporicidal activity of chemical disinfectants in the medical area. Test method and requirements (phase 2, step 1) [↑](#footnote-ref-28)
29. *Official Methods of Analysis of the AOAC International*, current edition, Sporicidal Activity of Disinfectants Test – Method II, Method 966.04 [↑](#footnote-ref-29)
30. European Standard EN 16615 Chemical disinfectants and antiseptics. Quantitative test method for the evaluation of bactericidal and yeasticidal activity on non-porous surfaces with mechanical action employing wipes in the medical area (4- field test). Test method and requirements (phase 2, step 2) [↑](#footnote-ref-30)
31. European Standard EN 13704 *Chemical disinfectants - Quantitative Suspension Test for the Evaluation of Sporicidal Activity of Chemical Disinfectants used in Food, Industrial, Domestic and Institutional areas - test method and requirements (phase 2, step 1)* [↑](#footnote-ref-31)